Automated Isoelectric Focusing Image Analysis for Improved Diagnosis and Classification of α1-Antitrypsin deficiency

Casey Vieni, Laura Eckelkamp, Jon Coker, Alex Popp, Melissa Snyder, David Murray, Patrick Vanderboom, Maria Willrich

α1-Antitrypsin (AAT) deficiency, results in increased activity of neutrophil elastase, with early onset emphysema and chronic obstructive pulmonary disease. AAT is measured by nephelometry and turbidimetry. Low concentrations of AAT require additional investigation. Some allele variants are associated with reduced AAT (S) and pathogenic, misfolding (Z) proteins compared to the normal M alleles. AAT phenotyping is an isoelectric focusing (IEF) method to characterize mobility patterns of the normal and variant alleles. The different  IEF patterns reflect amino acid substitutions and glycoforms. The objective of this proof of concept study was to evaluate if the gel manual-read out could be automated with the help of machine learning.

Methods: Twenty routine IEF gels (360 labelled instances) were interpreted by expert technologists. They were scanned and manual review was considered truth. We then trained machine learning models on 20 IEF gel images to classify AAT variant patterns into 4 classes: xxxxxxxx. The model was validated on 109 instances following preprocessing, including image registration, background subtraction with median denoising, and lane calling and realignment.

Results: A random forest (RF) model achieved 67% accuracy of the validation set compared to the training test set (macro-averaged precision=0.68, recall=0.57, F1-score=0.6). A convolutional neural network (CNN) slightly overfit the validation set with an accuracy of 0.78 (training set accuracy = 0.94). The models performed well on the most prevalent M variant (RF model: 91/104 correct, recall = 0.88, F1 = 0.75, CNN model: 94/102, recall = 0.92, F1 = 0.78) and the highly distinct MZ variant (RF model: 22/26 correct, precision = 0.85, recall = 0.85, F1 = 0.85; CNN model: 18/26, precision = 1.0, recall = 0.69, F1 = 0.82).  Limited number of samples led to poor performance in highly similar, but underrepresented variants, such as the MM1 variants (RF model: 12/38 correct, precision = 0.75, recall = 0.32, F1 = 0.44; CNN model: 14/38, precision=0.50, recall=0.37, F1 = 0.42), or rare variants (RF model: 12/28 correct, precision=0.55, recall = 0.43, F1=0.48; CNN model: 9/28, precision=0.69, recall = 0.32, F1 = 0.44). Singular value decomposition of gel images and feature weights of our model reveal classification was largely dominated by bands M2, M7, and M8.

Conclusions: The results of our pilot data suggest that increased data representation of AAT  variants may improve model performance, especially when coupled to the AAT concentration. This approach may improve the efficiency of expert human review through standardizing interpretation and streamlining classification of common and/or highly distinct AAT variants in the short-term, and combined to other methods can decrease the burden of this manual laboratory assay.